## Slide 1
WONIX: Training and Scanner Harmonization
Tyler Bradshaw
WON Spring Meeting
April 9, 2016

## Slide 2: WONIX objectives
Provide Wisconsin residents with access to advanced molecular imaging clinical trials
1. Increase availability of advanced molecular imaging agents
2. Improve imaging standards
3. Establish an informatics highway for sharing data
4. Create a minable data warehouse
5. Implement state-of-the-art clinical studies

## Slide 3: WONIX objectives
Provide Wisconsin residents with access to advanced molecular imaging clinical trials
1. Increase availability of advanced molecular imaging agents
2. Improve imaging standards
3. Establish an informatics highway for sharing data
4. Create a minable data warehouse
5. Implement state-of-the-art clinical studies
My role

## Slide 4: Quantitative PET imaging
Images are visual representations of numbers
2 3 4 5 5 4 3 2
1 4 5 7 7 6 5 2
2 5 7 9 9 8 5 2
2 5 8 9 9 8 5 2
2 5 7 9 8 6 3 2
2 3 5 6 5 4 2 1

## Slide 5: Quantitative PET imaging
In PET, if we’re careful, numbers can be used for comparisons
Max = 9
Max = 15
Comparison within a patient

## Slide 6: Quantitative PET imaging
In PET, if we’re careful, numbers can be used for comparisons
Max = 18
Comparison across patients
Max = 9

## Slide 7: Problems in quantitative PET imaging
Uncertainties, uncertainties, uncertainties!
Image values affected by uncertainties at almost every stage of imaging
Example: Errors up to 50%!

## Slide 8: Reducing uncertainties
Harmonize, harmonize, harmonize!
Harmonize patient preparation
Harmonize injection methods
Harmonize scanning parameters
Harmonize image processing
Harmonize image analysis

## Slide 9: Reducing uncertainties
Harmonize, harmonize, harmonize!
Harmonize patient preparation
Harmonize injection methods
Harmonize scanning parameters
Harmonize image processing
Harmonize image analysis
Site training

## Slide 10: WONIX – site training
WONIX
Administrator
Imaging Manager
Imaging QA Manager
Imaging Program Specialist
Site
Site Administrator
Study Coordinator(s)
Imaging Facility Rep(s)
1. WONIX Qualification Site Initiation Visit
2. WON Clinical Site Initiation Visit
WONIX
Administrator
Imaging Manager
Imaging Program Specialist
Site
Site Administrator
Study Coordinator(s)
Imaging Facility Rep(s)
Qualification Manual
Imaging Manual

## Slide 11: Reducing uncertainties
Harmonize, harmonize, harmonize!
Harmonize patient preparation
Harmonize injection methods
Harmonize scanning parameters
Harmonize image processing
Harmonize image analysis

## Slide 12: Goals of qualification
Ensure that sites can acquire high quality images
Ensure scanner is properly calibrated and maintained
Harmonize WONIX site’s PET/CT scanner
Qualification Manual

## Slide 13: Phantoms for quality assurance
Phantoms = objects we can scan to simulate patients
We know what the image numbers should be
WONIX qualification phantom scanned at each site
WONIX evaluates the image quality, scanner calibration
Phantom
Phantom PET image
Scanned at site
Evaluated by WONIX

## Slide 14: Goals of qualification
Ensure that sites can acquire high quality images
Ensure scanner is properly calibrated and maintained
Harmonize WONIX site’s PET/CT scanner
Qualification Manual

## Slide 15: PET scanner harmonization
SNMMI’s Clinical Trials Network (CTN) sent the same phantom to 170 sites, and collected and analyzed the images.
SUV = PET image value

## Slide 16: PET scanner harmonization
Early era
Mid-range
High end TOF
PRF
SNMMI’s Clinical Trials Network (CTN) sent the same phantom to 170 sites, and collected and analyzed the images.
LARGE DIFFERENCES BETWEEN SCANNERS!
SUV = PET image value

## Slide 17: PET scanner harmonization
WONIX method for scanner harmonization:

## Slide 18: PET scanner harmonization
Results SUV differences: 22% → 6% Noise differences: 24% → 0%
Melanoma patient scanned 3 months apart on 2 scanners

## Slide 19: Impact of scanner harmonization
A lung cancer patient had multiple PET/CT exams on the Discovery VCT (DVCT) and Discovery 710 (D710) scanners. SUVmax was measured over time with (red) and without (blue) harmonization.

## Slide 20: Harmonization workflow
Submit PET/CT images and form to WONIX UWCCC
Compare image quality to reference scanner
Harmonized?
Perform Qualification PET/CT scan
Reconstruct PET images
Send WONIX site the Harmonized Reconstruction Parameters
Send WONIX site PET/CT reconstruction request form
Y
N
WONIX site uses recon parameters for all subsequent imaging
Performed by WONIX site
Performed by WONIX UWCCC

## Slide 21: Reducing uncertainties
Harmonize, harmonize, harmonize!
Harmonize patient preparation
Harmonize injection methods
Harmonize scanning parameters
Harmonize image processing
Harmonize image analysis
Check scanner calibration
Evaluate image quality
Scanner harmonization
Image QA checks

## Slide 22: Reducing uncertainties
Harmonize, harmonize, harmonize!
Harmonize patient preparation
Harmonize injection methods
Harmonize scanning parameters
Harmonize image processing
Harmonize image analysis

## Slide 23: WONIX image analysis workflow
WONIX site
Image quality assurance
Image analysis
Database
WONIX QA and Analysis

## Slide 24: Quantitative PET imaging
In PET, if we’re careful, numbers can be used for comparisons
Max = 13
Goal: comparisons across SITES
Max = 9
WONIX Site 1
WONIX Site 2

## Slide 25: Thank you!
Informatics
Richard Bruce, Co-Chair
Sean Fain, Co-Chair
Imaging Clinical Trials
Robert Jeraj, Co-Chair
Glenn Liu, Co-Chair
Molecular Imaging Agents
Scott Perlman, Co-Chair
John Centanni, Co-Chair
Quantitative Imaging
Chihwa Song, Co-Chair
Tyler Bradshaw, Co-Chair
Working Groups
Data
Jonathan Boott, Co-Chair
Umberto Tachinardi, Co-Chair
WONIX Leadership
Robert Jeraj, Director
Scott Perlman, Co-Director
Dona Alberti, Administrator
WON
Ryan Mattison
Kari Wisinski
WON Representatives
WONIX Admin Support
Katie Kovacich
Mark McClintock
Angel Stampfli